MRI Implant Tips

Implants (Overview)

- Implant & Foreign Object

- Default 1.5T & 3T vs. MRI Lead Clearance

- IUD

- GI Clip

- VP Shunt

- Palate Expander

- Breast Tissue Expanders

- Baclofen Pump

- Deep Brain Stimulator

- Neurostimulators

Exam Time & Duration (Scheduling Windows)

| Specific Days | Time Needed | Duration | Exam |
| --- | --- | --- | --- |
| Monday–Friday | 8–4 PM | 1 hr | CSF Leak (Dr Delman) |
| Monday–Friday | 8–4 PM | 1.5 hr | CSF Leak Spine Survey |
| Monday–Friday | 8–4 PM | 1.5 hr | Diamox |
| Monday–Friday | 8–4 PM | Routine | MD Check Cases |
| Mon, Wed, Thu | 8–4 PM | 1 hr | Neurogram |
| Monday–Friday | 8–4 PM | 1 hr | Fetal |
| Mon–Thu | 9–2 PM | 1 hr | Pacemaker |
| Tue, Friday | 9–2 PM | 1.5 hr | Pediatric Cardiac |
| Monday–Friday | 8–4 PM | 1 hr | Pediatric Patient under 6 |

Implants: Default Clearance vs. Lead Clearance

Default for 1.5T & 3.0T

- Aneurysm coil

- Valve replacement / Annuloplasty band

- Dexcom glucose monitor — must be removed

- Breast clip (biopsy)

- Cardiac stent

- Hip/knee replacement (if not in the area of interest)

- All IUDs except Chinese Ring

- Orthopedic hardware (pin/screw/rod)

- Amplatzer occluder

- Scleral buckle

- Weight in eye (gold/platinum)

- Watchman device

Default for 1.5T

- IVC filter

- Harrington rod (if scanning spine)

- Shunt

- TIPS

- Palate expander

- Hardware in scanning region

- Braces / palate expander (for brain studies)

- Mechanical heart valve

Note: Any metal in the area of interest must remain at 1.5T.

Must Be Cleared by MRI Leads

- Aneurysm clip

- Breast tissue expander (MRI unsafe, except Motiva Flora—see below)

- Bullet / shrapnel

- Cochlear implant

- DBS (Deep Brain Stimulator)

- Defibrillator

- Loop recorder

- Pacemaker / pacing wires

- Penile implant

- Pump / infusion device

- Stimulator

- Sleep apnea device

Course of Action (By Implant / Object)

| Implant / Foreign Object | Imaging Needed | When | Field Strength | Time Slot | Notes |
| --- | --- | --- | --- | --- | --- |
| Loop recorder | Obtain card/info & notify leads | — | 1.5T/3.0T | — | — |
| Bullet | AP/Lateral X-ray | When X-ray available | 1.5T | — | See next slide for X-ray orders |
| Metal fragments in eye | CT Orbits | When CT available | 1.5T | — | CT orbits order required |
| Baclofen pump | KUB X-ray | Weekday 9–5 | 1.5T | — | Notify leads; rep present post-MRI |
| DBS | Shunt Series | Weekday (1 hr) | 1.5T | 1 hr | Notify leads; pt brings remote |
| Neurostimulators | Location of implant | Weekdays | 1.5T | — | Pt must bring remote; see workflow |
| Dexcom glucose monitor | None | Any | Any | — | Must be removed before entering MRI room |
| Abandoned leads | None | — | None | — | Not MRI safe — refer for different study |

Implant & Foreign Object — Documentation

| Device | Docs Needed | Comments |
| --- | --- | --- |
| Neurostimulator | Card; Remote & device fully charged | If batteries dead → need MD confirmation; if status can’t be checked → can’t be done |
| Cochlear implant | Card; some require MRI Nucleus KIT | Inform leads |
| Pacemaker/Defib | Card; Cardiology form (within 4–6 weeks) | Requires DOO on documentation; Chest X-ray for clearance |
| Deep brain stimulator | Card; Eligibility form | Can patient program device, or is a rep needed? |
| Abandoned lead | None | Not MRI conditional |
| Metal in orbits | X-ray or CT orbits (consult MD) | CT will require a script |
| BB / bullet | X-ray or check priors | MD clearance |
| Penile implant | Card | Coloplast up to 3T |

X-ray Clearance for Bullet Fragment

| Bullet Location | X-ray |
| --- | --- |
| Head | Skull series X-ray |
| Neck | Neck/Cervical |
| Chest | Chest AP/Lateral |
| Abdomen | AP/Lateral |
| Pelvis | AP / Hip series |
| Extremity | Appropriate extremity |

Schedulers: X-ray for MRI clearance does not need scripts. Inform patients our radiologist will review the X-ray prior to MRI; if inconclusive, a CT will be required. CT authorization may depend on insurance.

Intrauterine Contraceptive Device (IUD)

Unsafe: Chinese Ring (stainless steel) — MRI unsafe
Safe up to 3T: ParaGard, Mirena, Kyleena, Skyla, Copper IUD, Implanon implant

Reference: THE List (mrisafety.com)

Gastrointestinal Clips

MRI Unsafe

- Olympus Long Clip

- Olympus QuickClip2

- Olympus QuickClip Long

- Triclip Clip

These clips may remain ~9.4 days on average (can be longer). Prior to MRI, the physician should confirm via X-ray or endoscopy that no clip remains.

Action: Obtain KUB X-ray to confirm clip is no longer in body before proceeding.

Reference: Safety topic/article (mrisafety.com)

VP Shunts (Non-Programmable vs Programmable)

Non-Programmable VP Shunts

- Safe up to 1.5T

- Obtain information from patient

- If unable to obtain info → schedule Shunt Series X-ray prior to MRI

Programmable VP Shunts

- Field strength depends on vendor

- Obtain VP shunt card and setting

- STRATA: Needs Neurosurgery check post-MRI

- CERTAS: Does not need Neurosurgery check post-MRI

All programmable shunt patients must see Neurosurgery within 24 hr post-MRI for assessment.

3/8/2024 • Created by: Emily Ho

Examples (Programmable)

- Medtronic Strata shunt

- Codman Certas shunt

- Aesculap shunt
All listed are MRI-conditional up to 3T; require reassessment post-MRI unless noted.

Palate Expander

- Many orthodontics (metal braces, retainers, palatal expanders) can undergo MRI.

- Brain/neck MRI: expect significant artifact; heating discomfort can terminate exam.

- Recommendation: Remove at dentist prior to brain/neck MRI.

Breast Tissue Expander

- Temporary balloon device; later replaced with permanent saline/silicone implant.

- Most expanders use a magnetic injection port → MR safety issue.

- MRI cannot be performed until expander is removed.

- Exception (MRI-conditional): Motiva Flora Tissue Expander (contains RFID instead of magnet).

Reference: THE List (mrisafety.com)

Aneurysm Clip

- Pre-1985: some cerebral clips ferromagnetic

- 1993: single fatality from mis-identified clip (case report)

- Modern clips: generally MRI conditional up to 3T

- Must supply aneurysm clip information

Historically MRI-Unsafe Clips: Codman Vari-Angle, Downs Multi-positional, Drake Heifetz, Housepian, Kapp, McFadden, Pivot, Sundt-Kees Multi-Angle, Yasargil FD models.

Baclofen Pump

- Drug infusion pump; patient must bring pump ID card.

- Entering the magnet stalls the pump; infusion suspends during MRI and may have delayed recovery.

- Requires clinical programmer assessment within 24 hr post-MRI.

Scheduler checklist

- Confirm patient has programmer visit within 24 hr post-MRI.

- If not arranged: notify MRI leads to coordinate programmer on-site; otherwise exam cannot be done.

- KUB X-ray needed to confirm pump is not oriented 90° relative to MRI Z-axis.

Contrast Premedication Decision Tree

Regimen

- Prednisone 50 mg PO at 13 hr, 7 hr, and 1 hr before contrast injection

- Benadryl 50 mg PO 1 hr before contrast injection

Neurostimulator & DBS Workflow

1) Scheduling

- Obtain implant information

- Obtain Eligibility Form from referring provider (blank form available from MRI leads)

- Confirm IPG and lead placement

- Confirm MRI conditions/scan parameters

- Remind patient to bring remote fully charged and able to set MRI mode; if not, inform leads to arrange vendor rep on site

- Schedule Shunt Series X-ray prior to exam

- 1 hr slot, Mon–Fri, 9:00–5:00 PM, 1.5T only

  - If patient cannot set device to MRI mode/off → review with representative

2) Pre-MRI

- For patients without prior imaging: X-ray performed and cleared by MD (no abandoned leads)

- Eligibility form and scan parameters reviewed with technologist

- Patient must show leads that device is off / in MRI mode

3) Post-MRI

- Images reviewed with section MD

- Patient must show technologist/leads that device is back to normal

Deep Brain Stimulator — Why X-ray Matters

- Correct routing: leads connect to extensions routed on same side as IPG

- Incorrect routing: any extension routed to opposite side IPG → incorrect pathway

Shunt Series shows placement of leads and helps confirm routing before MRI.

Pacemaker & Cardiac Leads — Chest X-ray Importance

- Broken/abandoned leads risk lead heating, endocardial burns, induced arrhythmias, damage, and capture/sensing changes

- Mismatch of leads and generator (different manufacturers) is not MRI-conditional (often occurs when generator is replaced and original wires remain)

Scheduling must: obtain pacemaker card and info. Any broken/abandoned leads → Not MRI safe.

Pacemaker Workflow

1) Scheduling

- Obtain implant info & update pacemaker tracking form → MRI leadership clears the implant

- Obtain Cardiology Form (≥ 3 days before MRI)

- MRI leads notify vendor/EP team to be onsite

  - If Cardiology Form not received → reschedule and fill with next pacemaker patient

2) Pre-MRI

- If no prior X-ray, obtain Chest AP & Lateral; MD to clear

- If X-ray within last month, section MD reviews/clears prior

- Vendor/EP team programs pacemaker/ICD per cardiology recommendations

- RA Nurse or EP nurse monitors vitals throughout MRI

3) Post-MRI

- Images reviewed with section MD

- Vendor reprograms pacemaker/ICD to original settings

Pacemaker Clearance Form (Template Fields)

Patient Demographics

- Patient name

- DOB

- MRN

Attach Copy Of

- Pacemaker card

- ICD card

- Defibrillator card

Cardiologist/Electrophysiologist

- Name

- Phone

- Fax

Exam To Be Performed

Please follow this template when submitting pacemaker clearance requests to #AC_PacemakerTeam.